­Conclusions 20190131

* Contamination:
  + Comes most likely from DAPI in the machine reasons:
    - Unstained samples after stained are contaminated, but only with events with rather low DAPI fluorescence
    - Buffer sample is not contaminated when run after stained sample
  + My Conclusion: you have no contamination problem, because if you run two stained samples after each other, the DAPI has no effect since both have DAPI beforehand
  + BUT: you need to run the unstained first (and not tell Oluf)
* 2ml vs 15 ml and filter vs spin combinations:
  + Make a last trial experiment on Tuesday